## Session Tracking
| Tephra and Glass Analytical Sessions - most recent first | [13 empty] | Custom date format: | 0 |  |  | Only use results exported April 15, 2014 or later; Check for newest version.
| List all folders that contain an MDB file | Ignore TitaniQ or Titanium in Quartz or Pst or Pseudotachylite or Geol 101 or Chem lab or Petrology class folders | [15 empty] | Processing and Procedural Notes
| Year and Analytical Session | Session ID | Analytical method | Session ID column, etc. added | Blank Column J added (between beam and line) | Columns Added To "Normalized and Corrected" tab | Prelim Data, Analysts, Etc. Copied to "Normalized and Corrected" tab | Data Transferred from Preliminary Sheet | Sample notes & populations filled in | Na, K ( Si, Al) TDI included with processed results | Session ID column filled in | Analyst filled in | Catalog Number & Mount ID column filled in | Completed Processing Sheet | Checked and Compiled | Notes | Newest Data File | Procedural Changes | Results exported with updated processing | Checked that Drift Correction on Standdards is Enabled | Data in new blank-adjusting spreadsheet | New NKT Check (pre Apr 2013 only) | New BG Offset Formulas - Part 1 | New NKT, BHVO, Orth, Qtz reference (2 places) | New BG Adjust Formulas - Part 2 | New BG Adjust Formulas - Part 3 |  | Export Date | In Processing Sheet (revision date)
Compilation files:
| TEPHRA DATA COMPILATION CU1000-CU1145, Reference (2).xlsx | [12 empty] | All data through Fall 2016 inspected, finished for DB
| TEPHRA DATA COMPILATION CU1146-CU1303.xlsx | Many missing session ID, analyst, etc.; Some need sample info from catalog
TEPHRA DATA COMPILATION CU1304-CU1538.xlsx
TEPHRA DATA COMPILATION CU1539-16xx.xlsx
Year-Month-File(1st-2nd,etc.)-Day/Session in file (A, B, C)
Yr-Mo-File portion identifies the processing file, with letter added it identifies the day/session within the file
ID links unknown and reference data from same session
If no session reference data, can pull in representative method reference data instead
Letters for day/session because a session could span more than one date
2021
| 2021-07 (July 19-) Routine Tephra REFRESH | 2021-07-02-A (July 21), B (July 22), C (July 26), D (July 27) | CU routine tephra glass version 6.0
| 2021-07 (July 16-) Rhyolite Traces refresh and G-Probe |  | CU high-precision rhyolite glass with trace elements version 2.0
| ~May 15th during scheduled service visit |  |  | Moxtec detector back in Sp6; New spectrometer alignments (Will need to run fresh MAN calibrations)
| 2021-05 (May 10-) Routine Tephra | 2021-05-02-A (May 10), B (May 11), C (May 13) | CU routine tephra glass version 6.0
| 2021-05 (May 3-) Routine Tephra | 2021-05-01-A (May 3), B (May 5), C (May 6), D (May 7) | CU routine tephra glass version 6.0
| 2021-04 (April 21-) Routine Tephra updated MAN | 2021-04-02-A (Apr 21-22), B (Apr 23), C (Apr 27), D (Apr 28), E (Apr 29), F (Apr 30) | CU routine tephra glass version 6.0
| 2021-04 (April 13-) Routine Tephra updated MAN | 2021-04-01-A (Apr 13), B (Apr 14), C (Apr 15), D (Apr 16), E (Apr 19), F (Apr 20) | CU routine tephra glass version 6.0
| 2021-02 (Mar 29-) Routine Tephra updated MAN | 2021-03-05-A (Mar 30), B (Mar 31), C (Apr 6), D (Apr 7), E (Apr 8), F (Apr 9) | CU routine tephra glass version 6.0
| 2021-02 (Mar 23-) Routine Tephra updated MAN | 2021-03-04-A (Mar 24), B (Mar 25), C (Mar 26) | CU routine tephra glass version 6.0
| 2021-02 (Mar 15-) Routine Tephra updated MAN | 2021-03-03-A (Mar 15), B (Mar 16), C (Mar 17), D (Mar 18), E (Mar 19) | CU routine tephra glass version 6.0
| 2021-02 (Mar 8-) Routine Tephra updated MAN | 2021-03-02-A (Mar 8), B (Mar 9), C (Mar 10), D (Mar 11), E (Mar 12) | CU routine tephra glass version 6.0
| 2021-02 (Feb 11-) Routine Tephra updated MAN 1 | 2021-03-01-A (Mar 2), B (Mar 3), C (Mar 4), D (Mar 5) | CU routine tephra glass version 6.0
| 2021-02 (Feb 11-) Routine Tephra updated MAN 1 | 2021-02-02-A (Feb 11), B (Feb 12), C, D, E, F, G | CU routine tephra glass version 6.0
| February 3 - 14, 2021 |  |  | New routine tephra processing file, based on new Feb-2021 basalt traces processing file; Includes EarthChem column organization
Fresh MAN calibration for routine tephra analysis at 10 nA
| 2021-02 Basalt Traces - NEW with Zn Processed2 copy.xlsx | 2021-02-01-D (Feb 10) | CU high-precision mafic glass with trace elements version 2.2
| 2021-02-10 |  |  | New signal cable for Spectrometer 6 - fixed issue that emerged suddenly over previous weekend
| 2021-02 Basalt Traces - NEW with Zn Processed2 copy.xlsx | 2021-02-01-C (Feb 8) | CU high-precision mafic glass with trace elements version 2.2 | [12 empty] | Sp 6 - Bad Rb, possibly bad Sr due to signal cable issue
| 2021-02 Basalt Traces - NEW with Zn Processed2 copy.xlsx | 2021-02-01-A (Feb 3), B (Feb 5-7 major test run) | CU high-precision mafic glass with trace elements version 2.2
| February 3 - 14, 2021 |  |  | New basalt traces routine with Zn on spectrometer 5 RAP developed; Fresh MAN calibration for method; New data processing file with updated blank adjustments and Standards-based normaliztion adjustments
| 2021-02-01 |  |  | Rebuilt PHA for Spectrometers 3 and 4 - had elevated noise problem; issue may have been gradually emerging over previous few months
2020
| 2020-12 Balascio tephra (basalt traces, G-probe) | 2020-12-01-A (Dec 17), B (Dec 21), C (Dec 29)), D (Dec 30) | CU high-precision mafic glass with trace elements version 2.0
| 2020-12-23 |  |  | Re-install AM loaner P-10 detector in Sp6 due to noise on Moxtec window detector
| 2021-12-17 | Balascio, G-Probe runs | [13 empty] | Sp 6 - Bad Rb-Sr due to noisy detector
| 2020-12-10 |  |  | Install rebuilt PHA for Spectrometer 2 - fixes Sp2 noise and erratic counts issue
| 2021-11-17 | TitaniQ run for Petrology class
| 2020-11-06 |  |  | Repaired Moxtec window P-10 detector installed in Sp6
| November 04- 13, 2020 |  |  | New crystals installed from NC State; All spectrometers re-aligned
| 2020-08 (Aug 7-) Halls Cave crypto | 2020-08-01-A (Aug 25), B (Aug 26), C (Aug 27) | CU routine tephra glass version 6.0
| 2020-07 (July 9-26) tephra Fe-Ti oxides (note Sp2 band lost at end) | 2020-07-02-A (July 9,10,13), B (July 14-15), C (July 20-21), D (July 21-23), E (July 24-25), F (July 25-26)
| 2020-07 (July 3-8) tephra (Rhyolite traces method) | 2020-07-01-A (July 3-4), B (July 6), C (July 7-8) | CU high-precision rhyolite glass with trace elements version 2.0
| 2020-06 (June 23-29) tephra | 2020-06-01-A (Jun23), B(Jun24), C (Jun25), D (Jun26), E (Jun29) | CU routine tephra glass version 6.0
New Camera-Backup Computer
COVID SHUTDOWN
| 2020-03 (Mar 17-) Glass traces (Basalts method) | 2020-03-01-A (Mar 19) | CU high-precision mafic glass with trace elements version 2.0
| 2020-01 (Feb 16-) Glass traces - part 2 (Feb 26) Basalts method | 2020-02-04-A (Feb 26) | CU high-precision mafic glass with trace elements version 2.0 | New File
| 2020-01 (Feb 16-) Glass traces - part 1 (Feb 16-21) Basalts method | 2020-02-03-A (Feb 16), B(Feb 17), C (Feb 18), D (Feb 19), E (Feb 20), F (Feb 21) | CU high-precision mafic glass with trace elements version 2.0 | Mostly done - Check populations and outliers |  |  |  |  | (check blank correction and standards normalization)
| 2020-02 (Feb 11-) Routine tephra | 2020-02-02-A (Feb 11), B (Feb 25) | CU routine tephra glass version 6.0 | 1 More Day
| 2020-02 (Feb 10-) Glass Traces - BASALTS | 2020-02-01 b-A (Feb 12), B (Feb 13), C (Feb 14) | CU high-precision mafic glass with trace elements version 2.0 | Add Preliminary
| 2020-02 (Feb 10-) Glass Traces - RHYOLITES | 2020-02-01 - A (Feb 10), B (Feb 28), C (Mar 1-2) | CU high-precision rhyolite glass with trace elements version 2.0 | 2 More Days |  |  | Day 1 finished ONLY
| 2020-01 (Jan 7-Feb 4) Routine tephra - Part 2 (Jan 31-Feb4) | 2020-01-01-A (Jan7), B (Jan 22),C (Jan 27), D (Jan 28), E (Jan 30), F (Jan 31), G (Feb 4) | CU routine tephra glass version 6.0
| 2020-01 (Jan 7-Feb 4) Routine tephra - Part 1 (Jan 7-30) | 2020-01-01-A (Jan7), B (Jan 22),C (Jan 27), D (Jan 28), E (Jan 30), F (Jan 31), G (Feb 4) | CU routine tephra glass version 6.0 |  |  |  | Day 1 finished ONLY
2019
| 2019-12 (Dec 20-Jan 2) Glass traces | 2019-12-02-A (Dec 20), B (Dec 23),C (Dec 24), D (Jan 1) | CU high-precision rhyolite glass with trace elements version 2.0 | Finished |  |  |  |  | (check blank correction and standards normalization)
| 2019-12 (Dec 3-15) Routine tephra | 2019-12-01-A (Dec 3),B (Dec 4), C (Dec 6), D (Dec 10), E (Dec 11) | CU routine tephra glass version 6.0 | Finished |  |  |  |  | (check blank correction and standards normalization)
| 2019-11 (Nov12-20) Routine tephra | 2019-11-01-A (Nov 12), B (Nov 13), C (Nov 14), D (Nov 19), E (Nov 20) | CU routine tephra glass version 6.0 | Finished | (check blank correction and standards normalization)
| 2019-10 (Oct 30-Nov6) part 2 (Nov 5-6) | 2019-10-03-A (Oct30), B(Oct31), C(Nov1), D(Nov5), E(Nov6) | CU routine tephra glass version 6.0 | Finished | (check blank correction and standards normalization)
| 2019-10 (Oct 30-Nov6) part 1 (Oct30,31-Nov1) | 2019-10-03-A (Oct30), B(Oct31), C(Nov1), D(Nov5), E(Nov6) | CU routine tephra glass version 6.0 | Finished | (check blank correction and standards normalization)
| 2019-10 (Oct 7-23) Routine tephra - part 2 | 2019-10-02-A (Oct22), B(Oct23) | CU routine tephra glass version 6.0 | Finished | (check blank correction and standards normalization)
| 2019-10 (Oct 7-23) Routine tephra - part 1 | 2019-10-01-A (Oct8), B(Oct10), C(Oct11), D(Oct16) | CU routine tephra glass version 6.0 | Finished | (check blank correction and standards normalization)
| 2019-09 (Sept 6) Silicates-Traces new MAN - part 4 OXIDES (Nov 26, Dec 12,13,14) | 2019-09-05-A (Nov 26), B (Dec 12), C (Dec 13), D (Dec 14) | CU Fe-Ti oxides 1.0
| 2019-09 (Sept 6) Silicates-Traces new MAN - part 3 (Nov 11, 18) | 2019-09-04-A (Nov 11), B (Nov18) | CU high-precision rhyolite glass with trace elements version 2.0 | Finished
| 2019-09 (Sept 6) Silicates-Traces new MAN - part 2 (Oct28-29 G-probe) | 2019-09-03-A (Oct28-29) | CU high-precision mafic glass with trace elements version 2.0 | Finished
| 2019-09 (Sept 6) Silicates-Traces new MAN - part 1 (Oct17-18, 24-25) | 2019-09-02-A (Oct17), B(Oct18), C(Oct24), D(Oct25) | CU high-precision rhyolite glass with trace elements version 2.0
| 2019-08 (Aug 30) Tephra new MAN calibration after detector change | SKIP
New Probe Computer
| 2019-06 (June 3) G-probe 22 - 2019-05 G-probe processed (new template) | 2019-06-01-A (Jun3-4 am) | CU high-precision mafic glass with trace elements version 2.0 | [7 empty] | added session ID
| 2019-05 (May 8-16) Zieg Olivine-Plag - 2019-05 Zieg processed (new template) EDS | 2019-05-01b-A (May10), B(May15), C(May16-17), D(May22-23), E(May24-25), F(May28-29) | SKIP | [7 empty] | added session ID
| 2019-05 (May 8-16) Zieg Olivine-Plag - 2019-05 Zieg processed (new template) WDS | 2019-05-01a-A (May10), B(May15), C(May16-17), D(May22-23), E(May24-25), F(May28-29) | SKIP | [7 empty] | added session ID
| 2019-03 (March 26-29) Testing after Upgrade | SKIP | SKIP
| Routine tephra version 6 |  |  | Microprobe spectrometers reconfigured to 3 xenon + 3 p-10; Moxtec windows added; slits adjusted
| 2019-02 (Jan 24-25) INTAV trace glasses | 2019-01-02-A (Jan24-25) | CU high-precision mafic glass with trace elements version 1.3 | [7 empty] | added session ID
| 2019-01 (Jan 1-5) Silicates with traces - 2019-01 VT Silicates & Std Block Testing processed (WDS) | 2019-01-01b-A (Jan1), B(Jan2-4), C(Jan10,15-16), D(Jan17-20), E(Jan21-23) | CU high-precision mafic glass with trace elements version 1.3 | [7 empty] | added session ID
| 2019-01 (Jan 1-5) Silicates with traces - 2019-01 VT Silicates & Std Block Testing processed (EDS) | 2019-01-01a-A (Jan1), B(Jan2-4), C(Jan10,15-16), D(Jan17-20), E(Jan21-23) | CU high-precision mafic glass with trace elements version 1.3 | [7 empty] | added session ID
2018
| 2018-10 (Oct 5-20) Zieg ol, plg trace |  | SKIP
| 2018-10 (Oct 1-15) Tephra processed | 2018-10-01-A (Oct1), B(Oct2), C(Oct8), D(Oct15) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
2018-10 (Oct 1-12) Tephra
| 2018-09 (Sept 17-28) Tephra processed | 2018-09-02-A (Sept17), B (Sept24) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes | needs final check only
| 2018-09 (Sept 10-14) Tephra | 2018-09-01-A (Sept11), B (Sept12) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes | needs final check only
| 2018-08 (Aug 27-31, Sept 3-7) Tephra | 2018-08-01-A (Aug27), B (Sept4), C (Sept5), D (Sept7) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes | Yes
| 2018-07 (July 18-20-) Zieg ol,plg trace |  | SKIP
| 2018-07 (July 5-) Tephra | 2018-07-01-A (Jul5) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes | Yes
| 2018-06 (June 11-) Zieg ol,plg trace |  | SKIP
| 2018-06 (June 4) G-Probe, updated traces routine |  | CU high-precision mafic glass with trace elements version 1.2
| 2018-05 (May 29- June 1) Tephra | 2018-05-02-A (May29), B (May30), C (May31), D (Jun1) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes | Yes
| 2018-05 (May 21-25) Tephra | 2018-05-01-A (May21), B (May22), C (May23), D (May24), E (May25) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-04 (April 30-May 4) Tephra | 2018-04-05-A (Apr30), B (May1), C (May2), D (May3) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-04 (April 23-27) Tephra | 2018-04-04-A (Apr23), B (Apr25), C (Apr27) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes
| 2018-04 (April 16-20) Tephra | 2018-04-03-A (Apr16), B (Apr18), C (Apr20) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-04 (April 9-13) Tephra | 2018-04-02-A (Apr9), B (Apr11), C (Apr13) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-03 (Mar 26-30) & (April 2-6) - April part | 2018-04-01-A (Apr2), B (Apr4), C (Apr5), D (Apr6) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-03 (Mar 26-30) & (April 2-6) - March part | 2018-03-02-A (Mar26), B (Mar28), C (Mar29), D (Mar30) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | In Template | Yes | Yes | Yes
| 2018-03 (Mar 13-23) Tephra | 2018-03-01-A (Mar13), B (Mar15), C (Mar21), D (Mar23) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | partly | mostly | In Template | Yes | Yes | Yes
| 2018-01 (Jan 29-Feb 2) | 2018-01-03-A (Jan29), B (Feb1) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-01 (Jan 16-19) | 2018-01-02-A (Jan16), B (Jan17), C (Jan18), D (Jan19) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
| 2018-01 (Jan 09-13) | 2018-01-01-A (Jan10), B (Jan11), C (Jan12) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | mostly | In Template | Yes | Yes | Yes
CU routine tephra glass version 5.0
2017
| 2017-12 (Dec 21 on) Tephra | 2017-12-01-A (Nov 28-29), B (Nov 30) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes
| 2017-11 (Nov 28-Dec 1) Tephra and G-probe | 2017-11-02-A (Nov 28-29), B (Nov 30) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes
| 2017-11 (Nov 6-10) Tephra | 2017-11-01-A (Nov 7), B (Nov 9) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes |  | Yes
| 2017-10 (Oct 30-Nov 3) Tephra | 2017-10-05-A (Oct 30), B (Oct 31), C (Nov 1), D (Nov 2), E (Nov 3) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes |  | Yes
| 2017-10 (Oct 23-27) Tephra Cascades & K Adams | 2017-10-04-A (Oct 23), B (Oct 24), C (Oct 25), D (Oct 26), E (Oct 27) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes | Yes
| 2017-10 (Oct 16-20) Tephra Cascades | 2017-10-03-A (Oct 16), B (Oct 17), C (Oct 18), D (Oct 19), E (Oct 20) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes | Yes
| 2017-10 (Oct 9-13) Tephra Cascades | 2017-10-02-A (Oct 9), B (Oct 10), C (Oct 12), D (Oct 13) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes | Yes
| 2017-10 (Oct 2-6) Tephra Cascades | 2017-10-01-A (Oct 2), B (Oct 3), C (Oct 4), D (Oct 5) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes |  | Yes
| 2017-09 (Sept 25-30) Tephra (Cascades and Balascio) | 2017-09-02-A (Sept 25), B (Sept 26), C (Sept 28), D (Sept 29) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes | Yes | Yes | Yes
| 2017-09 (Sept 19-22) Tephra | 2017-09-01-A (Sept 19 - stds only), B (Sept 20), C (Sept 21), D (Sept 22) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template
| 2017-08 (Aug 29-30) Routine Tephra 14kV | 2017-08-01-A (Aug 30) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes | Yes | Yes | Yes |  | Updated MAN standards set and data | [6 empty] | Yes | Yes | Yes | Yes |  | 2016-07-05
| 2017-07 (July 10-14) Zieg thin sections (and G-Probe) - glass part | 2017-07-01-A,B,C | CU high-precision mafic glass with trace elements version 1.1 | [12 empty] | Reference glasses in new majors-traces routine
| 2017-07 (July 10-14) Zieg thin sections (and G-Probe) - mineral part | 2017-07-01-A,B,C | SKIP | [12 empty] | Reference glasses in new majors-traces routine
| 2017-07 (July 4-6) New Glass Traces (with G-Probe) |  | SKIP
Rebuilt PHAs installed; Spectrometers Aligned; New MAN Background standards
| 2017-05 (May 1-5) Tephra | 2017-05-01-A (May1), B (May2), C (May3), D (May4), E (May5) | CU routine tephra glass version 5.0
| 2017-04 (Apr 24-28) Tephra | 2017-03-03-A (Mar 21), B (Mar 22), C (Mar27) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template
| 2017-04 (Apr 18-21) Tephra | 2017-04-02-A (Apr19), B (Apr20), C (Apr21) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template
| 2017-04 (Apr 3-5) Tephra - Cascades | 2017-04-01-A (Apr19), B (Apr20), C (Apr21) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template
| 2017-03 (Mar 21-24) Tephra - Cascades | 2017-03-03-A (Mar 21), B (Mar 22), C (Mar27) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | Yes |  | Yes
| 2017-03 (Mar 16-17) Tephra - Cascades | 2017-03-02-A (Mar 16), B (Mar 17), C (Mar 18) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | Yes | Yes | Yes | Yes | Yes
| 2017-03 (Mar 10-12) Balascio tephra glass | 2017-03-01-A (Mar10), B (Mar11) | CU routine tephra glass version 5.0 | In Template | In Template | In Template | In Template | Yes | Yes | yes | Yes | Yes |  | Yes
| 2017-02 (Feb 27-Mar 1) Tephra - Cascades (Sp5 increased intensity) | 2017-02-01-A (Feb 28) | CU routine tephra glass version 5.0
| 2017-01 (Jan 22-25) Tephra - cascades - No useful data | SKIP | SKIP
| 2017-01 (Jan 4-6) General silicates majors & traces initial setup | SKIP | SKIP
| 2017-01 (Jan 3-5) Ti in Quartz initial setup |  | SKIP
2016
| 2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup 14kV |  | CU routine tephra glass version 5.0
| Routine tephra version 5 | Microprobe upgrade to 6 WDS spectrometers, Ba added |  | ADD interference corrections
| 2016-06 (June 27-July 4) Tephra | 2016-06-03-A (June29), B (June30), C (July1), D (July5) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | Yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-06 (June 27-July5) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-07-05
| 2016-06 (June 21-24) Tephra | 2016-06-02-A (June21), B (June22), C (June23), D (June24) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | Yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-06 (June 21-24) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-06-27
| 2016-06 (June 15-16) G-Probe glass Traces - No need to report | 2016-06-01-A (June15-16), B (June17-18) | SKIP | Yes | n/a | n/a | n/a | n/a | n/a | n/a | Yes | Yes |  | Yes | n/a |  | 2016-06 (June 15-16) RHYOLITE TRACES Sheet .xlsx | [5 empty] | n/a | n/a | n/a | n/a
| 2016-05 (May 23-27) Tephra | 2016-05-03-A (May23), B (May24), C (May25), D (May26) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-05 (May 23-27) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-05-27
| 2016-05 (May 10-13) Tephra | 2016-05-02-A (May10), B (May11), C (May12a), D (May12b) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-05 (May 10-13) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-05-13 | 2016-05-13
BEAM SPOT SIZE RECALIBRATED - previous spots smaller than reported (e.g. When 6 microns was reported, it was probably really 5 microns)
| 2016-05 (May 3-6) Tephra | 2016-05-01-A (May3), B (May4), C (May5), D (May6) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-05 (May 3-6) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-05-10 | 2016-05-27
| 2016-04 (Apr 18-22, 29) Tephra | 2016-04-02-A (Apr19), B (Apr20), C (Apr22), D (Apr29) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-04 (Apr 18-22) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-04-29 | 2016-05-27
| 2016-04 (Apr 8-14) Tephra (April 8) | 2016-04-01-A (Apr8) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-04 (Apr 8) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-04-19 | 2016-05-27
| 2016-03 (Mar 21-25) Tephra | 2016-03-02-A (Mar21), B (Mar22), C (Mar23), D (Mar25) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-03 (Mar 21-25) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-03-28 | 2016-05-27
| 2016-02 (Feb 29-Mar 4) Tephra | 2016-03-01-A (Mar2), B (Mar4), C (Mar11) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-02 (Feb 26) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-03-14 | 2016-05-27
| 2016-02 (Feb 15-19, 26) Tephra - Feb 26 part | 2016-02-04-A (Feb26) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-02 (Feb 15-19) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-02-29 | 2016-05-27
| 2016-02 (Feb 15-19, 26) Tephra - Feb 15-19 part | 2016-02-03-A (Feb15), B (Feb16), C (Feb17), D (Feb18), E (Feb19) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-02 (Feb 15-19) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-02-29 | 2016-05-27
| 2016-02 (Feb 8-12) Tephra | 2016-02-02-A (Feb8), B (Feb10), C (Feb11), D (Feb12) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | Yes |  | Yes | Yes |  | 2016-02 (Feb 8-12) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-02-15 | 2016-05-27
| 2016-02 (Feb 1-5) Tephra | 2016-02-01-A (Feb1), B (Feb2), C (Feb5) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | yes |  | Yes | Yes |  | 2016-02 (Feb 1-5) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-02-08 | 2016-03-01
| 2016-01 (Jan 25-29) Tephra | 2016-01-04-A (Jan27), B (Jan29) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | yes |  | Yes | Yes |  | 2016-01 (Jan 25-29) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-01-31 | 2016-02-29
| 2016-01 (Jan 18-22) Tephra | 2016-01-03-A (Jan18), B (Jan19), C (Jan20), D (Jan21) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | Yes | yes | Yes | Yes |  | Yes | YES | [7 empty] | Yes | Yes | Yes | Yes |  | 2016-01-22 | 2016-02-29
Filament controller/timer installed
| 2016-01 (Jan 12-15) Tephra | 2016-01-02-A (Jan13), B (Jan15) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | yes |  | Yes | Yes |  | 2016-01 (Jan 12-15) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-01-15 | 2016-02-26
| 2016-01 (Jan 5-8) Tephra, Noise Testing | 2016-01-01-A (Jan5), B (Jan6) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | Yes | yes |  | Yes | Yes |  | 2016-01 (Jan 5-8) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2016-01-25 | 2016-02-28
2015
| 2015-12 (Dec 21-23) Tephra | 2015-12-03-A (Dec21-22 until 3 am), B (Dec22 starting 10 am) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | yes | yes |  | Yes | Yes |  | 2015-12 (Dec 21-23) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-12-23
| 2015-12 (Dec 14-18) Tephra | 2015-12-02-A (Dec14), B (Dec15), C (Dec16), D (Dec17), E (Dec18) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | yes | yes |  | Yes | YES |  | 2015-12 (Dec 14-18) Tephra Processing.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-12-21
| 2015-12 (Dec 7-11) Tephra | 2015-12-01-A (Dec7), B (Dec8), C (Dec9), D (Dec11) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | yes | yes |  | Yes | YES |  | 2015-12 (Dec 7-11) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-12-14
| 2015-11,12 (Nov 30-Dec 4) Tephra | 2015-11-03-A (Nov30), B (Dec1), C (Dec2), D (Dec4) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes |  | 2015-11,12 (Nov 30-Dec 4) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-12-07
| 2015-11 (Nov 16-20) Tephra | 2015-11-02-A (Nov16), B (Nov17), C (Nov201) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes |  | yes | yes | yes |  | Yes | YES, but problem with Day 3 calibration |  | 2015-11 (Nov 16-20) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-11-20
| 2015-11 (Nov 9-13) Tephra | 2015-11-01-A (Nov9), B (Nov10), C (Nov11), D (Nov13) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? |  | yes | yes | yes |  | Yes | Yes |  | 2015-11 (Nov 9-13) tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-11-13 | 2015-11-16
| 2015-10 (Oct 27-30, Nov 6) Tephra | 2015-10-04-A (Oct27), B (Oct30), C (Nov6) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | yes | yes |  | Yes | Yes | [7 empty] | Yes | Yes | Yes | Yes |  | 2015-11-09
| 2015-10 (Oct 20-23) Tephra | 2015-10-03-A (Oct20) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | yes | yes | yes | yes |  | Yes | Yes | (Only one day has analyses) | [6 empty] | Yes | Yes | Yes | Yes |  | 2015-11-09
| 2015-10 (Oct 12-16) G-Probe (rhy traces) - Good Data | 2015-10-02-A (Oct17-19) | CU high-precision rhyolite glass with trace elements version 1.0 | Yes | n/a | n/a | yes | n/a | Yes | n/a | Yes | Yes |  | n/a | n/a |  | 2015-10 (Oct 16-19) RHYOLITE TRACES Sheet .xlsx | [5 empty] | n/a | n/a | n/a | n/a |  | 2015-10-20 | 2015-10-23
| 2015-10 (Oct 12-16) Tephra | 2015-10-01-A (Oct13), B (Oct14) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | Yes | yes | yes | yes |  | Yes | YES |  | 2015-10 (Oct 12-16) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-10-19 | 2015-11-01
2015-10 (Oct 5-9) CURRENT TESTING
| 2015-09 (Sept 28-Oct 2) Tephra | 2015-09-04-A (Sept28), B (Sept29) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | Yes | Yes | yes | yes | yes |  | Yes | YES |  | 2015-09 (Sept 28-Oct2) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-10-12 | 2015-11-01
| 2015-09 (Sept 21-25) Tephra | 2015-09-03-A (Sept21), B (Sept22), C (Sept23), D (Sept25) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | File fixed, add details for unknowns | yes | File fixed, add details for unknowns | File fixed, add details for unknowns |  | Yes | Yes |  | 2015-09 (Sept 21-25) Tephra Processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-10-12 | 2015-11-01
| 2015-09 (Sept 14-18) Tephra | 2015-09-02-A (Sept14), B (Sept15), C (Sept16), D (Sept18) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes |  | 2015-09 (Sept 14-18) OR-Gl Tephra Processed Data.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-09-21 | 2015-11-01
| 2015-09 (Sept 8-11) | 2015-09-01-A (Sept8), B (Sept9), C (Sept11) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes |  | 2015-09 (Sept 8-11) OR-Gl Tephra Processed Data.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-09-14 | 2015-11-01
| 2015-08 (Aug 25-28) Tephra | 2015-08-01-A (Aug25) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | Yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | 2015-08 (Aug25-28) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-10-12 | 2015-11-01
| 2015-07 (July 22-24) Tephra | 2015-07-01-A (Jul22), B (Jul28), C (Jul29), D (Jul30) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | [7 empty] | Yes | Yes | Yes | Yes |  | 2015-07-30 | 2015-08-06
Start of no analyst in preliminary data files - for older samples, use work schedules, project assignments, & note books instead
| 2015-05 (May 8,11-15) Tephra | 2015-05-01-A (May8), B (May11), C (May14) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes |  | 2015-05 (May 8,11-15) Tephra processed.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-05-14 | 2015-11-01
| 2015-04 (April27-May1) Tephra | 2015-04-04-A (Apr27), B (Apr28) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | Tephra Sheet OrGl 2015-04 (April27-28) Tephra.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-05-01 | 2015-05-01
| 2015-04 (April 21-24) Tephra | 2015-04-03-A (Apr21), B (Apr22), C (Apr24) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | Tephra Sheet OrGl 2015-04 (April 21-24) Tephra.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-04-28 | 2015-05-01
| 2015-04 (Apr 13-17) tephra | 2015-04-02-A (Apr13), B (Apr14), C (Apr16), D (Apr17), D (Apr18) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | YES | yes | yes | yes |  | Yes | Yes | some preliminary data not yet transferred | [6 empty] | Yes | Yes | Yes | Yes |  | 2015-04-20 | 2015-11-01
| 2015-04 (Apr6-10) tephra | 2015-04-01-A (Apr6), B (Apr7), C (Apr8), D (Apr10) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yea | yes |  | Yes | Yes | [7 empty] | Yes | Yes | Yes | Yes |  | 2015-04-12 | 2015-04-13
| 2015-03 (Mar27-28) Oxides-Sulfides at 14kV |  | Experimental ??
| 2015-03 (Mar26) G-Probe 14 |  | CU routine tephra glass version 4.0
| 2015-03 (Mar23-27) tephra | 2015-03-03-A (Mar23), B (Mar24), C (Mar25) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | Tephra Sheet OrGl 2015-03 (Mar23-27) tephra.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-03-25 | 2015-04-06
| 2015-03 (Mar2-6) tephra - Wed-Thurs part | 2015-03-02-A (Mar4), B (Mar5) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | Tephra Sheet OrGl 2015-03 (Mar2-6) tephra Wed-Thurs.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-03-05 | 2015-04-06
| 2015-03 (Mar2-6) tephra - Mon-Tues part | 2015-03-01-A (Mar2), B (Mar3) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | Tephra Sheet OrGl 2015-03 (Mar2-6) tephra Mon-Tues.xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-03-03 | 2015-03-05
| 2015-02 (Feb23-27) tephra | 2015-02-02-A (Feb23), B (Feb24), C (Feb25), D (Feb26), E (Feb27) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes | [6 empty] | n/a | Yes | Yes | Yes | Yes |  | 2015-03-02 | 2015-03-02
| 2015-02 (Feb15-20) tephra, new MAN data | 2015-02-01-A (Feb19), B (Feb20) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | Yes | yes |  | Yes | Yes |  | Tephra Sheet OrGl 2015-02 (Feb19-20).xlsx | [5 empty] | Yes | Yes | Yes | Yes |  | 2015-03-08 | 2015-11-01
2014
| 2014-11-12 (Nov-Dec) Zieg Thin Sections |  | SKIP for now
| 2014-11 (Nov 17-19) Tephra | 2014-11-02-A (Nov17), B (Nov18) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | [5 empty] | New Orth Glass spreadsheet (4) | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 2 days in run |  | 2014-12-22
| 2014-11 (Nov 9-13) Tephra | 2014-11-01-A (Nov10), B (Nov11), C (Nov12), D (Nov13), E (Nov14) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | [5 empty] | New Orth Glass spreadsheet (3) | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 5 days in run |  | 2014-11-14
| 2014-10 (Oct 20-24) Tephra | 2014-10-02-A (Oct20), B (Oct21), C (Oct22) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | MSH, Abbott (includes mafics) |  |  |  |  | New Orth Glass spreadsheet with linkeded ref values & Fe-Ca fix | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 3 days in run |  | 2014-10-24
| 2014-10 (Oct 9-10, 13-15) Tephra | 2014-10-01-A (Oct9), B (Oct10), C (Oct13), D (Oct14), E (Oct15) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | MSH, Abbott |  |  | NKT, BHVO for calibration; Blank correction |  | ON GLASS DATA SHEET | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 5 days in run |  | 2014-10-16
| 2014-09 (Sept 29-Oct 3) Tephra (Sept 30-31) | 2014-09-04-A (Sept29), B (Sept30) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | CU1532 - Burney ,CA (ref samples run on Tues) |  |  | NKT, BHVO for calibration; Blank correction |  | New Orth Glass spreadsheet with linkeded ref values & Fe-Ca fix | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 2 days in run |  | 2014-10-01
| 2014-09 (Sept 22-26) Tephra | 2014-09-03-A (Sept22), B (Sept23), C (Sept24), D (Sept25), E (Sept26) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | CU1524, 27, 33 - Burney, CA; CU1005-10 MSH; CU1099 Trego ref (bad?); CU1310 - Aegean |  |  | NKT, BHVO for calibration; Blank correction |  | New Orth Glass spreadsheet with linkeded ref values | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 5 days in run |  | 2014-09-29
| 2014-09 (Sept 15-19) Tephra and Geo 101 | 2014-09-02-A (Sept15), B (Sept16), C (Sept17), D (Sept18), E (Sept19) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | CU1308, 09, 20, 21 - Aegean; CU1001-05 MSH; CU1340 - McLaurin |  |  | NKT, BHVO for calibration; Blank correction |  | New Orth Glass spreadsheet with linkeded ref values | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 5 days in run |  | 2014-10-08
| 2014-09 (Sept 9-12) | 2014-09-01-A (Sept9), B (Sept10), C (Sept11), D (Sept12) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | MSH, Aegean tephras |  |  | NKT, BHVO for calibration; Blank correction | Yes | New Orth Glass spreadsheet with linkeded ref values & Fe-Ca fix | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 4 days in run |  | 2014-09-12
| 2014-07 (July 4) Rhyolite glass traces, GP-13 - MODIFIED, Good data | 2014-06-01-A (Jul4-5), B (Jul7), C (Jul8) | CU high-precision rhyolite glass with trace elements version 1.0 | Yes | Yes | Yes | yes | in compilation? | yes | n/a | yes | all S. Kuehn |  | Yes | Yes | G-probe & Newberry, Gl Pk reference |  | Modified rhyolite traces with Fe on WDS; Now only Al is on EDS | NKT, BHVO for calibration; Blank correction | Yes | Traces spreadsheet | n/a | n/a | Traces sheet
| 2014-06 (June 23-27, 30-July3) Tephra with Sp4 PET (Jun30-Jul3 part) | 2014-06-04-A (Jun30), B (Jul1), C (Jul2), D (Jul3) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | Aegean, MSH Set C |  |  | NKT, BHVO for calibration; Blank correction |  | X | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 4 days in run |  |  |  | Later half (June 30-July 3)
| 2014-06 (June 23-27, 30-July3) Tephra with Sp4 PET (Jun23-27 part) | 2014-06-03-A (Jun23), B (Jun25), C (Jun26), D (Jun27) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | Aegean, MSH Set C |  |  | NKT, BHVO for calibration; Blank correction |  | X | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 5 days in run |  |  |  | Earlier half (June 23-June 27)
| 2014-06 (June 18-20) Tephra with Sp4 PET | 2014-06-02-A (Jun18), B (Jun19), C (Jun20) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | yes | yes | yes |  | Yes | Yes | Aegean, MSH Set C |  |  | NKT, BHVO for calibration; Blank correction |  | X | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 3 days in run
| 2014-06 (June 10-13) Tephra with Sp4 PET | 2014-06-01-A (Jun11), B (Jun12), C (Jun13) | CU routine tephra glass version 4.0 | Yes | Yes | Yes | yes | in compilation? | yes | YES | yes | yes |  | Yes | Yes | Aegean, MSH Set C |  | Revised procedure with P, Cl on Sp4 with PET; Ca moved from LIF to PET; Orthoclase glass for blank correction; No more aggregate intensities | NKT, BHVO for calibration; Blank correction |  | X | n/a | Yes | Yes (EDS v2) | Weighted Formulas - NKT, Lip, BHVO, Orth glass | With Lipari Fe, Ca 4 days in run
| 2014-05 (May 27-29) Rhyolite glass traces - EXPERIMENTAL | 2014-05-02-A (May29-30) | SKIP | Yes |  |  |  | n/a | yes | n/a | yes | yes |  | Yes | Yes | Reference samples only |  | Testing new rhyolite traces procedure with MAN backgrounds; Sp4 changed to PET | n/a |  | Traces spreadsheet | n/a | n/a | Traces sheet
| Routine tephra version 4 | Sp4 crystal changed to PET - P and Cl moved to Sp4, Ca moved from LIF (3) to PET (1); Mg on RAP only | [10 empty] | No interference corrections
| 2014-05 (May13-14) Tephra - EDS Si, Al | 2014-05-01-A (May13), B (May14) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | yes | YES | Yes | yes |  | Yes | Yes | Aegean |  |  | NKT, BHVO for calibration |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  |  |  | Prepared spreadsheet which uses Orthoclase OR-1A instead of Orthoclase glass for blank
| 2014-04 (Apr 30-May 1, 7-8) Tephra - EDS Si, Al | 2014-04-05-A (Apr30), B (May1), C (May7), D (May8) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes | Carp Lake and comparative samples |  |  | NKT, BHVO for calibration |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run
| 2014-04 (Apr 23-24) Tephra - EDS Si, Al | 2014-04-04-A (Apr23), B (Apr24) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | needs populations | YES | yes | yes |  | Yes | Yes | Carp Lake and comparative samples |  |  | NKT, BHVO for calibration | OK | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run
| 2014-04 (Apr 15-17) Tephra - EDS Si, Al | 2014-04-03-A (Apr15), B (Apr16), C (Apr17) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes | Carp Lake |  |  | NKT, BHVO for calibration | Yes | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run
| 2014-04 (Apr 9-10) Tephra - EDS Si, Al | 2014-04-02-A (Apr9), B (Apr10) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes | Carp Lake |  | PET moved to Sp4 at end, but still using TAP for now | NKT, BHVO for calibration | Yes | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run
| 2014-04 (Apr 2-3,9) Tephra - EDS Si, Al | 2014-04-01-A (Apr2), B (Apr3) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes |  |  |  | NKT, BHVO for calibration | Yes | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run
| 2014-03 (Mar19-20,27) Tephra - EDS Si, Al | 2014-03-02-A (Mar19), B (Mar20), C (Mar27) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes |  |  | EDS with fixed element list | NKT, BHVO for calibration | Yes | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  |  |  | TDI rechecked | K TDI ranges from 0 to 3.5%; For 2wt% K2O, that is 0.06 wt% on a StDev of 0.3
| 2014-03 (Mar5-6) Tephra - EDS Si, Al | 2014-03-01-A (Mar5), B (Mar6) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes |  |  | No more WDS Si and Al, now EDS only | NKT, BHVO for calibration | Yes | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  |  |  | TDI rechecked
| 2014-02 (Feb 20, 27) Tephra | 2014-02-03-A (Feb20), B (Feb27) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | yes | YES | yes | yes |  | Yes | Yes |  |  |  |  | Y for WDS, Y for EDS | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 15-Apr, 6-Jul |  | Reprocessed both ways - EDS and WDS Si; Using EDS version only
| 2014-02 (Feb 16-17) G-probe 12 | 2014-02-03-A (Feb16-17) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | n/a | YES | yes | all S. Kuehn |  | Yes | Yes | Uses regular glass procedure |  | Note: Has fix to load EDS spectra with fle setup | NKT, BHVO for calibration | Y for WDS, Y for EDS | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 1 day in run |  | 2014-07-06 |  | Reprocessed both ways - EDS and WDS Si; Using EDS version only
| 2014-02 (Feb 12-13) Tephra | 2014-02-02-A (Feb12) | CU routine tephra glass version 3.4 | Yes | Yes | Yes | yes |  | yes | yes | yes | yes |  | Yes | Yes |  |  | Note: Has fix to load EDS spectra with fle setup | NKT, BHVO for calibration | Yes | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 1 day in run |  | 2014-07-06 |  | OK | EDS: Drift corr on, Aggregate off; Loaded Stds with no EDS all disabled
| 2014-01 (Jan 29-Feb 7) Tephra - Feb 5-6 part | 2014-02-01-A (Feb5), B (Feb6) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | yes | yes | yes | yes |  | Yes | Yes |  |  | Note: Before fix to load EDS spectra with fle setup | NKT, BHVO for calibration | Y for WDS, Y for EDS | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run |  | 2014-07-06 |  | Disabled all Stds Loaded with Setup (no EDS) to get processing to succeed; This requires using fewer MAN standards, only those analyzed in the specific run. | MAN standards: Ti (1, 17, 9, 90, 91), Fe (1, 17, 18, 69), Mn (1, 17, 18, 69), Mg (17, 18, 69, 91), Ca (17, 69, 91), Na (17, 69, 91), K (1, 17, 51, 69, 91), P (17, 18, 69, 91 curved), Cl (17, 18, 28, 69, 91)
| 2014-01 (Jan 29-Feb 7) Tephra - Jan 29 part | SKIP |  | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP | Jan 29 data is Junk - Bad Na |  | Note: Before fix to load EDS spectra with fle setup | n/a | n/a | SKIP - BAD DAY | SKIP | SKIP | SKIP | SKIP | SKIP |  | n/a |  | OMIT | Probably don't need aggregate Mg, unless have specific need for the higher precision on basaltic tephras; Otherwise, just use Spectro 2 RAP alone.
2013
All of these should start with 2013 - FIX
| Run 2013-11 (Nov 5-7) Tephra | 2014-11-1-A (Nov5), B (Nov6), C (Nov7) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | yes | yes | Yes | yes |  | Yes | Yes |  |  | Note: Before fix to load EDS spectra with fle setup | NKT, BHVO for calibration | Y | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-09-13 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | No Apatite in run with EDS Used NKT; Could pull P data from old spreadsheet
| Run 2013-10 (Oct 29-31) Tephra | 2014-10-3-A (Oct29), B (Oct30), C (Oct31) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  | Started, but not yet finished | yes | Yes | yes |  | Yes | Yes |  |  | Note: Before fix to load EDS spectra with fle setup | NKT, BHVO for calibration | Y | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-09-15 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | Has almost all standards
| Run 2013-10 (Oct 22-23) Tephra glasses | 2014-10-3-A (Oct22), B (Oct23) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  | Note: Before fix to load EDS spectra with fle setup | NKT, BHVO for calibration | Y | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-10-02 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | MAN standards: Ti (1, 17, 9, 90, 91), Fe (1, 17, 18, 69), Mn (1, 17, 18, 69), Mg (17, 18, 69, 91), Ca (17, 69, 91), Na (17, 69, 91), K (1, 17, 51, 69, 91), P (17, 18, 69, 91 curved), Cl (17, 18, 28, 69, 91) | No Apatite in run with EDS Used NKT; Could pull P data from old spreadsheet
| Run 2013-10 (Oct 15-17) Tephra glasses | 2014-10-3-A (Oct15), B (Oct16), C (Oct17) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  |  | NKT, BHVO for calibration | Y | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-10-02 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | MAN standards: Ti (1, 17, 9, 90, 91), Fe (1, 17, 18, 69), Mn (1, 17, 18, 69), Mg (17, 18, 69, 91), Ca (17, 69, 91), Na (17, 69, 91), K (1, 17, 51, 69, 91), P (17, 18, 69, 91 curved), Cl (17, 18, 28, 69, 91) | No Apatite in run with EDS Used NKT; Could pull P data from old spreadsheet
| Run 2013-10 (Oct 8-10) Tephra glasses | 2014-10-2-A (Oct8), B (Oct9), C (Oct10) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  |  | NKT, BHVO for calibration | Y | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-10-03 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | MAN standards: Ti (1, 17, 9, 90, 91), Fe (1, 17, 18, 69), Mn (1, 17, 18, 69), Mg (17, 18, 69, 91), Ca (17, 69, 91), Na (17, 69, 91), K (1, 17, 51, 69, 91), P (17, 18, 69, 91 curved), Cl (17, 18, 28, 69, 91) | No Apatite in run with EDS Used NKT; Could pull P data from old spreadsheet
| Run 2013-10 (Oct 1-3) Tephra glasses | 2014-10-1-A (Oct1), B (Oct3) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | [5 empty] | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-10-03 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | MAN standards: Ti (1, 17, 9, 90, 91), Fe (1, 17, 18, 69), Mn (1, 17, 18, 69), Mg (17, 18, 69, 91), Ca (17, 69, 91), Na (17, 69, 91), K (1, 17, 51, 69, 91), P (17, 18, 69, 91 curved), Cl (17, 18, 28, 69, 91)
| Run 2013-09 (Sept 17-19, 25-26) Tephra glasses | 2014-09-3-A (Sept17), B (Sept18), C (Sept19), D (Sept25), E (Sept26) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  |  | NKT, BHVO for calibration | Y | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 5 days in run |  | 2014-10-01 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | MAN standards: Ti (1, 17, 9, 90, 91), Fe (1, 17, 18, 69), Mn (1, 17, 18, 69), Mg (17, 18, 69, 91), Ca (17, 69, 91), Na (17, 69, 91), K (1, 17, 51, 69, 91), P (17, 18, 69, 91 curved), Cl (17, 18, 28, 69, 91)
| Run 2013-09 (Sept 10-12) Tephra glasses | 2014-09-2-A (Sept10), B (Sept11), C (Sept12) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | TDI was off for all unknowns - FIX |  |  |  |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-10-17 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | No Apatite in run with EDS Used NKT; Could pull P data from old spreadsheet
| Run 2013-09 (Sept 2-6) Tephra glasses | 2014-09-1-A (Sept2-3 until 3am), B (Sept3), C (Sept4), D (Sept5) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | [5 empty] | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run |  | 2014-10-18 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run
| Run 2013-07 (July 30-Aug 2) Alaska tephras | 2014-07-05-A (Jul31), B (Aug1), C (Aug2) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | Zander Unknowns on 31st, 1st, 2nd |  | Ca returned to Sp3 |  |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run |  | 2014-11-03 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, not MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | Missing MgO, MnO, Apatite
Sp3 crashed into mechanical limit 26-July; fixed July 28; New lead screw July 29
end Ca on PET
| Run 2013-07 (July) 15-19 folder (July 15-30 run) July 15,16, 17, 18 | 2013-07-02-A (Jul15), B (Jul16), C (Jul17), D (Jul18) | CU routine tephra glass version 3.4b (Ca on PET) | yes | Yes | Yes | yes |  |  | Yes | Yes | yes |  | Yes | Yes | Process in three files: July 15, 16, 17, 18 (4 days) |  | Ca on Sp1 PET temporarily; problem with bearings near Ca position on Sp3 |  |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run |  |  |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | No MgO
| Run 2013-07 (July) 15-19 folder (July 15-30 run) July 22, 23, 24, 25, 26 | 2013-07-03-A (Jul22), B (Jul23), C (Jul24), D (Jul25), E (Jul26) | CU routine tephra glass version 3.4b (Ca on PET) | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | Process in three files: July 22, 23, 24, 25, 26 (5 days) |  | Ca on Sp1 PET temporarily; problem with bearings near Ca position on Sp3 |  |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 5 days in run |  |  |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | No MgO
| Run 2013-07 (July) 15-19 folder (July 15-30 run) July 29,30 | 2013-07-04-A (Jul29), B (Jul30) | CU routine tephra glass version 3.4b (Ca on PET) | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | Process in three files: July 29, 30 (2 days) |  | Ca on Sp1 PET temporarily; problem with bearings near Ca position on Sp3 |  |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-10-24 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Clear and redo MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | No MgO
start Ca on PET
| Run 2013-07 (July) 10-13 Sum Lk FF Site tephras | 2013-07-01-A (Jul10), B (Jul11) skip 12-13th | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | SP3 bearing problem on Ca may start on April 2 |  |  |  |  | X (OR-1a&Qtz version) | n/a | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-10-30 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run
| Run 2013-06 (June) G-probe 30nA glass | 2014-06-01-A (Jun12) still in old format data file | CU tephra glass version 3.4, 30 nA |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | Check date format |  |  | SKIP | SKIP | [5 empty] | SKIP
30 nA | 2013-05 (May) 18-24 General Silicates |  | ADD LATER
| Run 2013-04 (Apr) 25 (AK & Carp) - April 25-May 2 | 2014-04-02-A (Apr25), B (Apr30), C (May2) omit NKT grid | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | Yes | Yes | yes |  | Yes | Yes | SP3 bearing problem on Ca may start on April 2 |  | Watch out for bad data on NKT before here - Snowflake crystals in NKT discovered during this run |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run |  | 2014-11-03 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - Run based on copied template MDB from Feb 25
| Run 2013-04 (Apr) 9,11 (Alaska & Carp Lake) | 2014-04-01-A (Apr9), B (Apr10), C (Apr11), D (Apr16) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | Yes | Yes | yes |  | Yes | Yes | SP3 bearing problem on Ca may start on April 2 |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 4 days in run |  | 2014-11-03 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - Run based on copied template MDB from Feb 25
| Run 2013-03,04 (Mar 26-28, Apr 2,4 (Carp Lake) | 2014-03-01-A (Mar26), B (Mar27), C (Apr2), D (Apr4), E (Mar28) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | Yes | Yes | yes |  | Yes | Yes | SP3 bearing problem may have affected Mn here too |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 5 days in run |  | 2014-11-03 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - Run based on copied template MDB from Feb 25
| Run 2013-02 (Feb) 25-27 (Carp Lake) | 2014-02-02-A (Feb25), B (Feb26) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 1 day in run |  | 2014-11-04 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - This run has standards used in subsequent runs
| Run 2013-02 (Feb) 18 (Carp Lake) | SKIP | CU routine tephra glass version 3.4 | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP | Incomplete Run - had to replace Sp1 leadscrew |  |  |  |  | SKIP | SKIP | SKIP
| 2013-02 (Feb) 12 (Carp Lake) | 2014-02-01-A (Feb12), B (Feb13) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-11-17 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - Run based on copied template MDB from Jan; Some deleted entries due to intermittent counter-timer errors
| Run 2013-01 (Jan) 30,Feb7 - (Carp Lake) | 2014-01-05-A (Jan31), B (Feb7) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | LEAD SCREW ISSUES Jan 31-Feb 5 |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 3 days in run |  | 2014-11-05 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards Run Jan 30; Manually fixed bad blank entry in EDS table; Some deleted entries due to intermittent counter-timer errors
| Run 2013-01 (Jan) 22-23 - (Carp Lake) | 2014-01-04-A (Jan22), B (Jan23) | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-11-19 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - Run based on copied template MDB from Jan 15
| Run 2013-01 (Jan) 14-15,16 - (Carp Lake): 16&21 | 2014-01-03-A (Jan16) omit 21st | CU routine tephra glass version 3.4 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes | Large drift on Sp3 Ca and Fe on Jan 21st |  | Mn changed to MAN background; Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-11-22 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run - Run based on copied template MDB from Jan 15
| Routine tephra version 3.4 | Mn changed to MAN background |  | No interference corrections
| Run 2013-01 (Jan) 14-15,16 - (Carp Lake): 14-15 | 2014-01-02-A (Jan14), B (Jan15) | CU routine tephra glass version 3.3 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 2 days in run |  | 2014-11-22 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run
| Run 2013-01 (Jan) 7 - (Carp Lake): 7-8 | 2014-01-01-A (Jan13) | CU routine tephra glass version 3.3 | yes | Yes | Yes | yes |  |  | yes | Yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT |  |  | X (OR-1a&Qtz version) | New NKT Check | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines (156-7), 1 before NKT end set (181/2) | With Lipari Fe 1 day in run |  | 2014-11-22 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Lipari, MnO as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards in this run
2012
| Run 2012-12 (Dec) 5 - Alaska glass | 2012-12-01-A (Dec 6 early am), B (Dec 6 7 am -Dec 7 2:30 am), C (Dec 7 7:39am -Dec 8 1:57 am), D (Dec 8 8:36a, - Dec 9 6:58 am), E (Dec 9 10:59am - Dec 10) | CU routine tephra glass version 3.3 | yes | Yes | Yes | yes |  |  | Yes | yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT | Exported twice - TAP for Mg, RAP for Mg |  | X (OR-1a&Qtz version) | New NKT Check Formatting | Yes | Yes, Bg adjust & Std Eval, ver 3c | Delete lines on Day 1: 2 orth std lines, 1 before NKT end set | With Lipari Fe 5 days in run |  | 2014-11-17 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP/RAP Mg. 6) NKT, Orthoclase, MnO as standards. 7) EDS Si, Al on Orthoclase. 8) Set TDI. 9) Set H2O by difference
Room Air Conditioner Installed November 19-20, 2012 - Now have stable temperatures all year; Earlier runs may have issues with thermal peak drift
| Run 2012-11 (Nov) 13 - GP-9 intermediate glass | 2012-11-01-A | CU tephra glass version 3.3, 15-40 | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | skip | SKIP | all D. Kuehn |  | SKIP | SKIP | No tephra samples in run |  | P changed to MAN background; Watch out for bad data on NKT | 40 nA, First Run with MAN for P; Only Mn with measured BG |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP
| Routine tephra version 3.3 | P changed to MAN background
| Run 2012-11 (Nov) 7 - tephra glass with EDS, MAN - Don't use | SKIP | SKIP | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP | Peak reproducibility problem on SP1 Si |  | Watch out for bad data on NKT |  |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP
| Run 2012-10 (Oct) 29-Nov 2 - tephra glass with EDS, MAN (Carp Lake) - Exported in two parts: Oct 29-31; Nov 1-2 | 2012-10-02-A (Oct 29), B (Oct 30), C (Oct 31), D (Nov 1), E (Nov 2) | CU routine tephra glass version 3.2 | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Yes | Yes | Intermittent missing counts for Mg on RAP; Thermal drift, filament out Oct 30; No EDS spectra (standby) first set on Nov 1, but do have standards over lunch + end; Sp1 peak repro issue shows up |  | Ti changed to MAN background; Watch out for bad data on NKT | P, Mn have measured BG |  |  | New NKT Check Formatting | Yes | Yes, Bg adjust & Std Eval, ver 3c | Yes; Lipari only for measured BGs | With Lipari Fe 5 days in run |  | 2015-01-14 |  | 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) NKT, Orthoclase, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | Used copy of previous MDB to avoid problem loading EDS spectra with setup
| Routine tephra version 3.2 | Ti, Mg, Cl changed to MAN background; Only Mn, P have measured backgrounds
| Run 2012-10 (Oct) 22 - tephra glass with EDS, MAN - Don't use | SKIP | SKIP | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP | [5 empty] | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP
| Run 2012-10 (Oct) 8 - tephra glass with EDS, MAN (Carp Lake) | 2012-10-02-A (Oct 8), B (Oct 10) | CU routine tephra glass version 3.1 | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Yes | Yes | Use Oct 8, 10; Skip Oct 9 |  | Watch out for bad data on NKT | P, Ti, Mg, Mn have measured BG |  |  | New NKT Check Formatting | Yes | Yes, Bg adjust & Std Eval, ver 3c | Yes; Lipari only for measured BGs | With Lipari Fe 2 days in run |  | 2015-01-14 |  | 1) Disable standards loaded with setup 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) BHVO, Sodalite, Spess as standards. Not using NKT for calibration. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards Run; Manually fixed bad blank entry in EDS table; Some deleted entries due to intermittent counter-timer errors
| Run 2012-10 (Oct) 3,4 - tephra glass with EDS, MAN (Carp Lake) | 2012-10-01-A (Oct 4) omit 5th | CU routine tephra glass version 3.1 | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Yes | Yes |  |  | Watch out for bad data on NKT | P, Ti, Mg, Mn have measured BG | First evaluation of NKT as possible calibration standard |  | New NKT Check Formatting | Yes | Yes, Bg adjust & Std Eval, ver 3c | Yes; Lipari only for measured BGs | With Lipari Fe 2 days in run |  | 2015-01-15 |  | 1) Disable all Stds Loaded with Setup. 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 4) Recheck MAN assignments. 5) Disable Si, Al, and TAP Mg. 6) NKT, BHVO, MnO as standards. 7) EDS Si, Al on BHVO. 8) Set TDI. 9) Set H2O by difference | all standards Run; Manually fixed bad blank entry in EDS table; Some deleted entries due to intermittent counter-timer errors
| Run 2012-09 (Sept) 24-Oct1 tephra glass MAN testing - Don't use | SKIP | SKIP | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP |  |  | First attempt at MAN backgrounds |  |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP
| Routine tephra version 3.1 | MAN Backgrounds begin - Fe, K, Na, Ca by MAN ; Begin 14 kV for routine glass
| Run 2012-09 (Sept) 11-14 tephra glass with EDS | 2012-09-01-A (Sept12), B (Sept13-14) | CU routine tephra glass version 2.0 | yes | Yes | Yes | yes |  |  | YES | yes | yes |  |  | Yes |  |  |  | All backgrounds measured |  |  | n/a | n/a | Yes, Bg adjust & Std Eval, ver 3c | n/a | n/a |  | 2015-01-15 |  | 1) Disable standards loaded with setup 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) Use traditional standards - BHVO, Lipari were run as unknowns 7) EDS Si, Al on Orthoclase. 8) Set TDI. 9) Set H2O by difference
Run 2012-06 (June) 13-14 PHA, Dead Time
| Run 2012-06 (June) 7, 17 tephra glass with EDS (June 21-23) | 2012-06-01-A (Jun 21), B (June 22), C (June 23) | CU routine tephra glass version 2.0 | yes | Yes | Yes | yes |  |  | YES | yes | yes |  |  | Yes |  |  |  | All backgrounds measured |  |  | n/a | n/a | Yes, Bg adjust & Std Eval, ver 3c | n/a | n/a |  | 2015-01-15 |  | 1) Disable standards loaded with setup 2) Turn off aggregate intensities. 3) Turn on drift correction on standards. 5) Disable Si, Al, and TAP Mg. 6) Use traditional standards - BHVO, Lipari were run as unknowns 7) EDS Si, Al on Orthoclase. 8) Set TDI. 9) Set H2O by difference
| Routine tephra version 2 | EDS+WDS Analyses begin
| Run 2012-04 (Apr) 18-19 WTB SPI - Don't use; Special MAS prescribed method | SKIP | SKIP
| Run 2012-04 (Apr) 6-7 glasses on glass block | 2012-04-02-A | CU routine tephra glass version 1.0, 30 nA
| Run 2012-04 (Apr) 4-5 tephra for UPS | 2012-04-01-A (Apr 5) | CU routine tephra glass version 1.0, 15 nA | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Re-Do in new processing sheet, but with blank correction off | Yes | Re-exported data from PFE 2016-Sept-27 | Run 2012-04 (Apr) 4-5 tephra NEW SHEET.xlsx | NO TDI% IN ORIGINAL EXCEL FILE | All backgrounds measured |  | Use Original Spreadsheets; Lipari & BHVO only for adjustments; Adjust all elements with newer formulas | n/a | n/a | New BHVO reference | n/a | n/a
| Run 2012-03 (Mar) 28-30 WTB traces - Trace only and initial versions of Basalt traces, Rhyolite traces methods in this folder | SKIP | SKIP | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP |  |  |  | All backgrounds measured |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP
| Run 2012-03 (Mar) 25-26 WTB majors | 2012-03-01-A | CU routine tephra glass version 1.0, 6 nA | SKIP |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP |  | SKIP | SKIP |  |  |  | All backgrounds measured |  | SKIP | SKIP | SKIP | SKIP | SKIP | SKIP
| Run 2012-01 (Jan) 30-21 | 2012-01-02-A (Jan 30), B (Jan 31 up to 5:21 pm), C (Jan 31 9:27 pm to Feb 1 7:24 am), D (Feb 1 9 am on) | CU routine tephra glass version 1.0, 6 nA | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Re-Do in new processing sheet, but with blank correction off | Yes | Re-exported data from PFE 2016-Sept-27, added water-by-difference | 2012-01 (Jan 30-31) 5 micron MSH NEW SHEET.xlsx | NO TDI% IN ORIGINAL EXCEL FILE | All backgrounds measured |  | Use Original Spreadsheets; Lipari & BHVO only for adjustments; Adjust all elements with newer formulas | n/a | n/a | New BHVO reference | n/a | n/a
| Run 2012-01 (Jan) 26-27 (stds) - Thermal drift, Don't use | SKIP
| Run 2012-01 (Jan) 15,16 | 2012-01-01-A (Jan 16), B (Jan 21), C (Jan 24) | CU routine tephra glass version 1.0, 6 nA | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Re-Do in new processing sheet, but with blank correction off | Yes | Re-exported data from PFE 2016-Sept-27 | 2012-01 (Jan) 15-24 (MSH Set C) NEW SHEET.xlsx | NO TIME STAMP OR TDI% IN ORIGINAL EXCEL FILE | All backgrounds measured |  | SKIP | n/a | n/a | New BHVO reference | n/a | n/a
2010-2011
| Run 2011-12 (Dec) 9,10 (stds) - Don't use | SKIP | SKIP
| Run 2011-11 (Nov) 29,30 (G-probe, Stds) | 2011-11-01-A | CU routine tephra glass version 1.0, 15 nA
| Run 2011-11 (Nov) 21,22 (stds) - Don't use | SKIP | SKIP
| Run 2011-10,11 (Oct 5 - Nov 16) | 2011-10-01-A (Oct 20-21 early am), B (Oct 21), C (Oct 27) | CU routine tephra glass version 1.0, 15 nA | yes | Yes | Yes | yes |  |  | YES | yes | yes |  | Re-Do in new processing sheet, but with blank correction off | Yes | Re-exported data from PFE 2016-Sept-27; Required use of "Concord-original" standard file due to change in standard numbers | 2011-10 (Oct5-Nov16) NEW SHEET..xlsx | NO TIME STAMP OR TDI% IN ORIGINAL EXCEL FILE | All backgrounds measured |  | Use Original Spreadsheets; Lipari & BHVO only for adjustments; Adjust all elements with newer formulas | n/a | n/a | New BHVO reference | n/a | n/a
| Run 2010-2011 (pre Oct-2011) - INTAV Run |  | CU routine tephra glass version 1.0, 15 nA
| Routine tephra version 1 | Initial data collection and testing begins

## Probe Methods
January 2011 |  |  | CU routine tephra glass version 1.0, 15 nA | [8 empty] | Full method description file finished
INTAV session
Last (Current) On and Off Peak Count Times:
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
BGDS: | LIN | LIN | LIN | LIN | HIGH | HIGH | LIN | LIN | LIN | LIN | LIN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 2
ONTIM: | 10 | 40 | 21 | 28 | 12 | 42 | 100 | 50 | 45 | 40 | 42 | 0
HITIM: | 3 | 15 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
LOTIM: | 3 | 15 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
7 TDI intervals | 15 kV | 15 nA
October 2011 |  |  | CU routine tephra glass version 1.0, 15 nA | [8 empty] | Full method description file finished
Initial routine tephra | [11 empty] | Run 2011-10,11 (Oct5-Nov16) Volcanic Glass.MDB
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
BGDS: | LIN | LIN | LIN | LIN | HIGH | HIGH | LIN | LIN | LIN | LIN | LIN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 2
ONTIM: | 10 | 28 | 21 | 28 | 12 | 42 | 100 | 50 | 45 | 40 | 42 | 0
HITIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
LOTIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
STDS: | 4 | 19 | 16 | 28 | 6 | 6 | 28 | 28 | 14 | 28 | 6 | 28
7 TDI intervals | 15 kV | 15 nA |  | Same as above, but shorter count time for P
Initial routine tephra ELEM:: January 2012; Initial routine tephra Cl ka: CU routine tephra glass version 1.0, 6 nA; Run 2011-10,11 (Oct5-Nov16) Volcanic Glass.MDB Mg ka: Full method description file finished | Detection limits for 6 nA from this one
Routine tephra - MSH samples | [11 empty] | Glass 2012-01-15 MSH.MDB
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
BGDS: | LIN | LIN | LIN | LIN | HIGH | HIGH | LIN | LIN | LIN | LIN | LIN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 2
ONTIM: | 10 | 28 | 21 | 28 | 12 | 42 | 100 | 50 | 45 | 40 | 42 | 0
HITIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
LOTIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
7 TDI intervals | 15 kV | 6 nA |  | Same as above, but lower beam current
Routine tephra - MSH samples ELEM:: April 2012; Routine tephra - MSH samples Cl ka: CU routine tephra glass version 1.0, 15 nA; Glass 2012-01-15 MSH.MDB Mg ka: Full method description file finished | Detection limits for 15 nA from this one
Routine tephra - Tepper | [11 empty] | Glass 2012-04-04.MDB
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
BGDS: | LIN | LIN | LIN | LIN | HIGH | HIGH | LIN | LIN | LIN | LIN | LIN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 2
ONTIM: | 10 | 28 | 21 | 28 | 12 | 30 | 110 | 50 | 45 | 40 | 30 | 0
HITIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
LOTIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5 | 0
6 TDI intervals | 15 kV | 15 nA |  | Same as above, but changed Al, Mg, and Na count times and TDI from 7 to 6
Routine tephra - Tepper ELEM:: April 2012; Routine tephra - Tepper Cl ka: CU routine tephra glass version 1.0, 30 nA; Glass 2012-04-04.MDB Mg ka: Full method description file finished | Detection limits for 30 nA from this one
Glass block
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
BGDS: | LIN | LIN | LIN | LIN | HIGH | HIGH | LIN | LIN | LIN | LIN | LIN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 2
ONTIM: | 10 | 28 | 21 | 28 | 12 | 30 | 110 | 50 | 45 | 40 | 30
HITIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5
LOTIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5
6 TDI intervals | 15 kV | 30 nA |  | Same as above, but 30 nA
EDS installed
June 2012 |  |  | CU routine tephra glass version 2.0 | [8 empty] | Full method description file finished
Glass block ELEM:: tephra glass with EDS; Glass block Mg ka: Tephra glass with EDS - testing.MDB | Detection limits from this file
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
BGDS: | LIN | LIN | LIN | LIN | HIGH | HIGH | LIN | LIN | LIN | LIN | LIN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 2
ONTIM: | 10 | 28 | 21 | 28 | 12 | 30 | 110 | 50 | 45 | 40 | 30
HITIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5
LOTIM: | 3 | 13 | 6 | 10 | 4 | 5 | 43 | 10 | 22 | 10 | 5
| 15 kV | 10 nA |  | Si and Al by EDS begins |  |  | EDS 130 kcps, 2 cm distance (in log book for June 19)
October 2012 |  |  | CU routine tephra glass version 3.1 | [8 empty] | Full method description file finished
Tephra glass with EDS - testing.MDB Mg ka: Run 2012-10 (Oct) 4 - glass with EDS, MAN - Copy for EDS | Detection limits from this file
ELEM: | K ka | P ka | Cl ka | Ti ka | Si ka | Al ka | Mg ka | Ca ka | Mn ka | Fe ka | Na ka | Mg ka
BGD: | MAN | OFF | MAN | OFF | MAN | MAN | OFF | MAN | OFF | MAN | MAN | OFF
BGDS: | MAN | LIN | MAN | LIN | MAN | MAN | LIN | MAN | LIN | MAN | MAN | LIN
SPEC: | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 4 | 4
CRYST: | PET | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | TAP | TAP
ORDER: | 2 | 5 | 4 | 3 | 1 | 1 | 2 | 1 | 3 | 2 | 1
ONTIM: | 10 | 18 | 15 | 24 | 12 | 30 | 84 | 50 | 38 | 40 | 30
HITIM: | ---- | 9 | ---- | 10 | ---- | ---- | 31 | ---- | 18 | ---- | ----
LOTIM: | ---- | 9 | ---- | 10 | ---- | ---- | 31 | ---- | 18 | ---- | ----
| 14 kV | 10 nA
October 2012 |  |  | CU routine tephra glass version 3.2 | [8 empty] | Full method description file finished
Run 2012-10 (Oct) 4 - glass with EDS, MAN - Copy for EDS Mg ka: Copy of Run 2012-10 (Oct) 29 - tephra5 - new standards for WDS.MDB | Detection limits from this file
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | OFF | MAN | MAN | MAN | MAN | OFF | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | LIN | MAN | MAN | MAN | MAN | LIN | MAN | MAN
SPEC: | 1 | 1 | 2 | 3 | 3 | 2 | 3 | 4 | 1 | 1 | 1 | 4
CRYST: | PET | PET | RAP | LiF | LiF | RAP | LiF | TAP | PET | PET | PET | TAP
ORDER: | 1 | 3 | 1 | 2 | 3 | 2 | 1 | 1 | 2 | 5 | 4
ONTIM: | 12 | 25 | 29 | 40 | 28 | 140 | 50 | 30 | 20 | 18 | 16
HITIM: | ---- | ---- | ---- | ---- | 12 | ---- | ---- | ---- | ---- | 9 | ----
LOTIM: | ---- | ---- | ---- | ---- | 12 | ---- | ---- | ---- | ---- | 9 | ----
| 14 kV | 10 nA
November 2012 |  |  | CU tephra glass version 3.3, 40 | [8 empty] | Full method description file finished
ELEM:: Gprobe; Copy of Run 2012-10 (Oct) 29 - tephra5 - new standards for WDS.MDB Mg ka: Run 2012-11 (Nov) 13 - G-probe glass.MDB | Detection limits for 40 nA from this one
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | OFF | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | LIN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 2 | 1 | 2 | 3 | 3 | 2 | 3 | 4 | 1 | 1 | 1 | 4
CRYST: | RAP | PET | RAP | LiF | LiF | RAP | LiF | TAP | PET | PET | PET | TAP
ORDER: | 1 | 2 | 2 | 2 | 3 | 3 | 1 | 1 | 1 | 4 | 3 | 2
ONTIM: | 15 | 40 | 29 | 40 | 28 | 114 | 50 | 30 | 30 | 40 | 25 | 140
HITIM: | ---- | ---- | ---- | ---- | 12 | ---- | ---- | ---- | ---- | ---- | ---- | ----
LOTIM: | ---- | ---- | ---- | ---- | 12 | ---- | ---- | ---- | ---- | ---- | ---- | ----
STDS: | 17 | 22 | 52 | 22 | 51 | 22 | 22 | 52 | 18 | 23 | 52 | 22
6 TDI intervals | 15 kV | 40 nA |  |  |  | EDS 9 cm distance (in log book)
December 2012 |  |  | CU routine tephra glass version 3.3 | [8 empty] | Full method description file finished
Gprobe ELEM:: Alaska; Run 2012-11 (Nov) 13 - G-probe glass.MDB Mg ka: Run 2012-12 (Dec) 5 - Alaska glass (2).MDB | Detection limits from this file
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | OFF | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | LIN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 2 | 1 | 2 | 3 | 3 | 2 | 3 | 4 | 1 | 1 | 1 | 4
CRYST: | RAP | PET | RAP | LiF | LiF | RAP | LiF | TAP | PET | PET | PET | TAP
ORDER: | 1 | 2 | 2 | 2 | 3 | 3 | 1 | 1 | 1 | 4 | 3 | 2
ONTIM: | 15 | 40 | 29 | 40 | 28 | 114 | 50 | 30 | 30 | 40 | 25 | 0
HITIM: | ---- | ---- | ---- | ---- | 12 | ---- | ---- | ---- | ---- | ---- | ---- | ----
LOTIM: | ---- | ---- | ---- | ---- | 12 | ---- | ---- | ---- | ---- | ---- | ---- | ----
STDS: | 18 | 22 | 9 | 22 | 51 | 22 | 22 | 52 | 18 | 23 | 52 | 22
| 14 kV | 10 nA
Checked
January 2013 |  |  | CU routine tephra glass version 3.4 | [8 empty] | Full method description file finished
Alaska ELEM:: Carp - 16th+; Run 2012-12 (Dec) 5 - Alaska glass (2).MDB Mg ka: Run 2013-01 (Jan) 16,21 - Carp Lake MAN for Mn.MDB | Detection limits from this file
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 2 | 3 | 3 | 2 | 3 | 4 | 1 | 1 | 1 | 4
CRYST: | PET | PET | RAP | LiF | LiF | RAP | LiF | TAP | PET | PET | PET | TAP
ORDER: | 1 | 3 | 1 | 2 | 3 | 2 | 1 | 1 | 2 | 5 | 4 | 2
ONTIM: | 15 | 28 | 29 | 55 | 38 | 140 | 55 | 30 | 24 | 28 | 22 | 0
| 14 kV | 10 nA |  | All backgrounds MAN
May 2013 |  |  | CU routine tephra glass version 3.4, 30 nA | [8 empty] | Full method description file finished
Carp - 16th+ ELEM:: G-probe; Run 2013-01 (Jan) 16,21 - Carp Lake MAN for Mn.MDB Mg ka: Run 2013-11 G-probe glass.MDB | Detection limits from this file
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 2 | 3 | 3 | 2 | 3 | 4 | 1 | 1 | 1 | 4
CRYST: | PET | PET | RAP | LiF | LiF | RAP | LiF | TAP | PET | PET | PET | TAP
ORDER: | 1 | 3 | 1 | 2 | 3 | 2 | 1 | 1 | 2 | 5 | 4
ONTIM: | 15 | 28 | 29 | 55 | 38 | 140 | 55 | 30 | 24 | 28 | 22
| 14 kV | 30 nA
Checked
July 2013 | CU routine tephra glass version 3.4b (Ca on Sp1 instead of 3 temporarily) | [10 empty] | Full method description file finished
Run 2013-11 G-probe glass.MDB Mg ka: Run 2013-07 (July) 15-19 Sum Lk (2).MDB | Detection limits from this file
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 2 | 3 | 3 | 2 | 1 | 4 | 1 | 1 | 1 | 4
CRYST: | PET | PET | RAP | LiF | LiF | RAP | PET | TAP | PET | PET | PET | TAP
ORDER: | 1 | 3 | 1 | 1 | 2 | 2 | 4 | 1 | 2 | 6 | 5 | 2
ONTIM: | 15 | 28 | 29 | 105 | 100 | 174 | 20 | 30 | 24 | 28 | 22 | 0
| 14 kV | 10 nA
May 2013 |  |  | General Silicates version 1.0
Silicates (Zieg first run)
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | Mg ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 4 | 1 | 2 | 3 | 3 | 2 | 1 | 4 | 1 | 4
CRYST: | TAP | PET | RAP | LiF | LiF | RAP | PET | TAP | PET | TAP
ORDER: | 2 | 3 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 3
ONTIM: | 25 | 35 | 40 | 40 | 60 | 60 | 30 | 40 | 30 | 0
same file, with Ni
ELEM: | Si ka | Ti ka | Al ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | Mg ka | Ni ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | OFF
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | LIN
SPEC: | 4 | 1 | 2 | 3 | 3 | 2 | 1 | 4 | 1 | 4 | 3
CRYST: | TAP | PET | RAP | LiF | LiF | RAP | PET | TAP | PET | TAP | LiF
ORDER: | 2 | 3 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 3 | 3
ONTIM: | 25 | 35 | 40 | 15 | 25 | 60 | 30 | 25 | 30 | 0 | 28
HITIM: | ---- | ---- | ---- | ---- | ---- | ---- | ---- | ---- | ---- | ---- | 12
LOTIM: | ---- | ---- | ---- | ---- | ---- | ---- | ---- | ---- | ---- | ---- | 12
simpler for feldspars
ELEM: | Si ka | Al ka | Fe ka | Ca ka | Na ka | K ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 4 | 2 | 3 | 1 | 4 | 1
CRYST: | TAP | RAP | LiF | PET | TAP | PET
ORDER: | 2 | 1 | 1 | 2 | 1 | 1
ONTIM: | 14 | 50 | 50 | 20 | 25 | 25
| 20 kv | 30 nA
Spectrometer 4 changes to PET for routine use
May 2014 |  |  | Experimental, skip
same file, with Ni ELEM:: 2014-05 (May 27-29) Rhyolite glass traces | 2014-05 (May) Rhyolite traces test Data.MDB
same file, with Ni ELEM:: July 2014; same file, with Ni Si ka: CU high-precision rhyolite glass with trace elements version 1.0 | Full method description file finished
same file, with Ni ELEM:: 2014-07 (July 4) Rhyolite glass traces, GP-13 | 2014-07 (July 4-8) Rhy traces, GP-13.MDB | Detection limits from this file
ELEM: | K ka | Ca ka | Cl ka | P ka | Na ka | Mg ka | Fe ka | Mn ka | Ti ka | Ba la | Si ka | Sr la | Zr la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 3 | 4 | 4 | 4
CRYST: | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | LiF | PET | PET | PET
ORDER: | 1 | 2 | 3 | 4 | 1 | 2 | 1 | 2 | 3 | 4 | 1 | 2 | 3
ONTIM: | 80 | 25 | 87 | 150 | 10 | 342 | 40 | 74 | 90 | 150 | 20 | 178 | 160
| 20 kV | 40 nA |  | EDS 15 s
EDS 8 cm distance (in log book)
June 2014 |  |  | CU routine tephra glass version 4.0 | [8 empty] | Full method description file finished
2014-06 (June 10-13) Tephra with Sp4 PET | [11 empty] | 2014-06 (June 10-11) Tephra Sp4 PET 2.MDB | Detection limits from this file
ELEM: | Ti ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 3 | 3 | 2 | 1 | 2 | 1 | 4 | 4
CRYST: | PET | LiF | LiF | RAP | PET | RAP | PET | PET | PET
ORDER: | 3 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 2
ONTIM: | 62 | 100 | 70 | 140 | 50 | 30 | 50 | 100 | 65
2014-06 (June 10-13) Tephra with Sp4 PET Ti ka: 14 kV; 2014-06 (June 10-13) Tephra with Sp4 PET Fe ka: 10 nA; 2014-06 (June 10-13) Tephra with Sp4 PET Mg ka: All backgrounds MAN; 2014-06 (June 10-13) Tephra with Sp4 PET K ka: Si and Al by EDS only | Spectro 4 now PET, Mg only on RAP
2014-06 (June 10-13) Tephra with Sp4 PET ELEM:: November 2014; 2014-06 (June 10-13) Tephra with Sp4 PET Mn ka: CU routine tephra glass version 4.0 |  |  | Full method description file finished
2014-06 (June 10-13) Tephra with Sp4 PET ELEM:: 2014-11 (Nov 9-13) Tephra |  |  | 2014-11 (Nov 9-13) Tephra.MDB | Previously processed detection limits in this folder
ELEM: | Ti ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 3 | 3 | 2 | 1 | 2 | 1 | 4 | 4
CRYST: | PET | LiF | LiF | RAP | PET | RAP | PET | PET | PET
ORDER: | 3 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 2
ONTIM: | 62 | 100 | 70 | 140 | 50 | 30 | 50 | 100 | 65
| 14 kV | 10 nA |  | All backgrounds MAN |  |  | EDS 8 s
2014-11 (Nov 9-13) Tephra ELEM:: June 2016; 2014-11 (Nov 9-13) Tephra Mn ka: CU routine tephra glass version 4.0 |  |  | Full method description file finished
2016-06 (June 27-July 4) Tephra
ELEM: | Ti ka | Fe ka | Mn ka | Mg ka | Ca ka | Na ka | K ka | P ka | Cl ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 3 | 3 | 2 | 1 | 2 | 1 | 4 | 4
CRYST: | PET | LiF | LiF | RAP | PET | RAP | PET | PET | PET |  | Stable method used for ~2 years
ORDER: | 3 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 2
ONTIM: | 62 | 100 | 70 | 140 | 50 | 30 | 50 | 100 | 65
| 14 kV | 10 nA |  | All backgrounds MAN |  |  | EDS 8 s
2015-03 (Mar27-28) Oxides-Sulfides at 14kV
2016-06 (June 27-July 4) Tephra ELEM:: October 2015; 2016-06 (June 27-July 4) Tephra Ti ka: CU high-precision rhyolite glass with trace elements version 1.0 |  |  | Full method description file finished |  | Same as July 2014 traces run
2016-06 (June 27-July 4) Tephra ELEM:: 2015-10 (Oct 12-16) G-Probe (rhy traces) - Good Data |  |  | 2015-10 (Oct16-17) Rhy traces, GP-15.MDB
ELEM: | K ka | Ca ka | Cl ka | P ka | Na ka | Mg ka | Fe ka | Mn ka | Ti ka | Ba la | Si ka | Sr la | Zr la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 3 | 4 | 4 | 4
CRYST: | PET | PET | PET | PET | RAP | RAP | LiF | LiF | LiF | LiF | PET | PET | PET
ORDER: | 1 | 2 | 3 | 4 | 1 | 2 | 1 | 2 | 3 | 4 | 1 | 2 | 3
ONTIM: | 80 | 25 | 87 | 150 | 10 | 342 | 40 | 74 | 90 | 150 | 20 | 178 | 160
| 20 kV | 40 nA |  | EDS 15 s
Change to 6 WDS Spectrometers
December 2016 |  | CU routine tephra glass version 5.0 | [9 empty] | Full method description file finished
2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup | [11 empty] | New 6WDS Glass Setup and GP17 (2 EDS at 2cm).MDB
ELEM: | k ka | ca ka | cl ka | fe ka | mn ka | na ka | mg ka | ba la | p ka | ti ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 4 | 5 | 6
CRYST: | PET | PET | PET | LiF | LiF | RAP | RAP | LiF | PET | PET
ORDER: | 1 | 2 | 3 | 1 | 2 | 1 | 2 | 1 | 1 | 1
ONTIM: | 45 | 45 | 60 | 90 | 70 | 30 | 130 | 165 | 165 | 165
STDS: | 99 | 114 | 172 | 114 | 145 | 114 | 114 | 135 | 134 | 114
| 20 kV | 40 nA |  | EDS 15 s
February 2017 |  | CU routine tephra glass version 5.0
2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup ELEM:: 2017-02 (Feb 27-Mar 1) Tephra - Cascades (Sp5 increased intensity) |  | 2017-02 (Feb 27-30) Tephra - Cascades.MDB
STDS: | 99 | 114 | 172 | 114 | 145 | 114 | 114 | 89 | 134 | 114
January 2018 |  | CU routine tephra glass version 5.0
[12 empty] | 2018-01 (Jan 16-19) tephra.MDB | Detection limits from this file
Full method description file finished
January 2017
2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup ELEM:: 2017-01 (Jan 4-6) General silicates majors & traces initial setup |  | 2017-01 (Jan 4-6) General silicates 6WDS.MDB
No Data Acquired
2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup ELEM:: July 2017; 2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup k ka: CU high-precision mafic glass with trace elements version 1.1 |  | Full method description file finished
2017-07 (July 4-6) New Glass Traces (with G-Probe)
2016-12 (Dec 16-20) GP17 and New 6WDS Glass Setup ELEM:: 2017-07 (July 10-14) Zieg thin sections (and G-Probe) |  | Zieg TS based off glass traces 2 - Copy for GP with EDS Si.MDB | Detection limits from this file (July 21, including Rb)
ELEM: | k ka | ca ka | cl ka | s ka | fe ka | co ka | ni ka | na ka | al ka | mg ka | ba la | mn ka | si ka | sr la | p ka | zr la | ti ka | v ka | cr ka | rb la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 3 | 3 | 3 | 4 | 4 | 5 | 5 | 5 | 5 | 6 | 6 | 6 | 6
CRYST: | PET | PET | PET | PET | LiF | LiF | LiF | RAP | RAP | RAP | LiF | LiF | PET | PET | PET | PET | PET | PET | PET | PET
ORDER: | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4
ONTIM: | 45 | 20 | 60 | 105 | 24 | 118 | 112 | 60 | 25 | 160 | 190 | 67 | 10 | 100 | 50 | 84 | 40 | 40 | 40 | 100
STDS: | 99 | 114 | 172 | 135 | 114 | 149 | 150 | 114 | 113 | 114 | 167 | 146 | 113 | 179 | 134 | 181 | 114 | 209 | 144 | 202
| 15 kV | 40 nA |  | EDS 24 s |  | Si and Al both EDS and WDS available
WDS Al starts on July 19
Rb starts on July 21
June 2018 | CU high-precision mafic glass with trace elements version 1.2 | [10 empty] | Full method description file finished
2018-06 (June 4) G-Probe, updated traces routine | [11 empty] | Glass Traces updated (V on LIF), GP 20, 21.MDB | Detection limits from this file
ELEM: | k ka | ca ka | cl ka | s ka | fe ka | co ka | ni ka | na ka | al ka | mg ka | ba la | v ka | mn ka | si ka | sr la | p ka | zr la | ti ka | cr ka | rb la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 3 | 3 | 3 | 4 | 4 | 4 | 5 | 5 | 5 | 5 | 6 | 6 | 6
CRYST: | PET | PET | PET | PET | LiF | LiF | LiF | RAP | RAP | RAP | LiF | LiF | LiF | PET | PET | PET | PET | PET | PET | PET
ORDER: | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 4 | 1 | 2 | 3
ONTIM: | 40 | 25 | 70 | 116 | 24 | 125 | 125 | 25 | 25 | 215 | 130 | 100 | 40 | 10 | 110 | 50 | 95 | 40 | 85 | 123
STDS: | 99 | 114 | 172 | 135 | 114 | 149 | 150 | 114 | 113 | 114 | 167 | 210 | 146 | 123 | 179 | 134 | 181 | 114 | 144 | 202
| 15 kV | 40 nA |  | EDS 12 s |  | Si and Al both EDS and WDS available
Move V from PET to LIF; Adjust count times
January 2019 | CU high-precision mafic glass with trace elements version 1.3 | [10 empty] | Full method description file finished
2019-01 (Jan 1-5) Silicates with traces | [11 empty] | 2019-01 VT Silicates with traces | Detection limits from this file
ELEM: | k ka | ca ka | cl ka | s ka | fe ka | co ka | ni ka | zn ka | na ka | al ka | mg ka | ba la | v ka | mn ka | si ka | sr la | p ka | zr la | ti ka | cr ka | rb la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 3 | 3 | 3 | 4 | 4 | 4 | 5 | 5 | 5 | 5 | 6 | 6 | 6
CRYST: | PET | PET | PET | PET | LiF | LiF | LiF | LiF | RAP | RAP | RAP | LiF | LiF | LiF | PET | PET | PET | PET | PET | PET | PET
ORDER: | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 4 | 1 | 2 | 3
ONTIM: | 40 | 24 | 100 | 120 | 24 | 93 | 90 | 90 | 25 | 25 | 246 | 130 | 115 | 55 | 10 | 114 | 60 | 114 | 40 | 104 | 135
STDS: | 99 | 114 | 172 | 135 | 114 | 149 | 150 | 426 | 114 | 114 | 114 | 167 | 210 | 146 | 114 | 179 | 134 | 181 | 114 | 144 | 202
| 15 kV | 40 nA |  | EDS 12 s |  | Si and Al both EDS and WDS available
Add Zn; Adjust count times
January 2019 | CU high-precision mafic glass with trace elements version 1.3
2019-02 (Jan 24-25) INTAV trace glasses | [11 empty] | 2019-02 (Jan 24-25) INTAV trace glasses.MDB
same as above
Microprobe spectrometers reconfigured to 3 xenon + 3 p-10; Moxtec windows added; slits adjusted
January 2018 |  | CU routine tephra glass version 6.0 | [9 empty] | Full method description file finished
2019-10 (Oct 7-23) Routine tephra | [11 empty] | 2019-10 (Oct 7-) Routine tephra.MDB
ELEM: | Ba la | Fe ka | Mn ka | P ka | Cl ka | K ka | Ca ka | Na ka | Mg ka | Ti ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 2 | 2 | 3 | 3 | 4 | 4 | 5 | 5 | 6
CRYST: | LiF | LiF | LiF | PET | PET | PET | PET | RAP | RAP | PET
ORDER: | 1 | 1 | 2 | 1 | 2 | 1 | 2 | 1 | 2 | 1
ONTIM: | 180 | 100 | 69 | 141 | 25 | 70 | 101 | 30 | 140 | 180
| 14 kV | 10 nA |  | EDS 8 s
2019-10 (Oct 7-23) Routine tephra ELEM:: June 2019; 2019-10 (Oct 7-23) Routine tephra Ba la: CU high-precision mafic glass with trace elements version 2.0 |  | Full method description file finished
2019-10 (Oct 7-23) Routine tephra ELEM:: 2019-06 (June 3) G-probe 22 - 2019-05 G-probe processed (new template) |  | 2019-05 (May 8-16) Zieg WDS with G-Probe added.MDB
ELEM: | Ba la | V ka | Mn ka | Fe ka | Co ka | Ni ka | Zn ka | Ca ka | P ka | Zr la | S ka | Cl ka | K ka | Ti ka | Cr ka | Na ka | Al ka | Mg ka | Si ka | Rb la | Sr la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 4 | 4 | 4 | 5 | 5 | 5 | 6 | 6 | 6
CRYST: | LiF | LiF | LiF | LiF | LiF | LiF | LiF | PET | PET | PET | PET | PET | PET | PET | PET | RAP | RAP | RAP | PET | PET | PET
ORDER: | 1 | 2 | 3 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3
ONTIM: | 135 | 115 | 65 | 25 | 95 | 95 | 95 | 20 | 55 | 100 | 80 | 50 | 40 | 100 | 175 | 25 | 25 | 265 | 25 | 160 | 130
| 15 kV | 40 nA |  | EDS 12 s |  | Si and Al both EDS and WDS available
Multiple methods in the same file
Fall 2019 | CU high-precision rhyolite glass with trace elements version 2.0
2019-09 (Sept 6) Silicates-Traces new MAN - part 3 (Nov 11, 18)
2019-09 (Sept 6) Silicates-Traces new MAN | [11 empty] | 2019-09-06 Silicates-Traces update MAN calibration.MDB | 2019-09 (Sept 6) Silicates-Traces new MAN - part 1 (Oct17-18, 24-25)
ELEM: | Ba la | Fe ka | Mn ka | Ca ka | P ka | Zr la | Cl ka | K ka | Ti ka | Na ka | Mg ka | Si ka | Rb la | Sr la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 2 | 2 | 3 | 3 | 3 | 3 | 4 | 4 | 5 | 5 | 6 | 6 | 6
CRYST: | LiF | LiF | LiF | PET | PET | PET | PET | PET | PET | RAP | RAP | PET | PET | PET
ORDER: | 1 | 1 | 2 | 1 | 2 | 3 | 4 | 1 | 2 | 1 | 2 | 1 | 2 | 3
ONTIM: | 338 | 100 | 226 | 35 | 80 | 117 | 62 | 50 | 275 | 25 | 300 | 17 | 170 | 130
| 15 kV | 40 nA |  | EDS 30 s |  | Si and Al both EDS and WDS available
Fall 2019 | CU high-precision mafic glass with trace elements version 2.0 (with timings adjusted) | [10 empty] | Full method description file finished
Detection limits from this file
2019-09 (Sept 6) Silicates-Traces new MAN | [11 empty] | 2019-09-06 Silicates-Traces update MAN calibration.MDB | 2019-09 (Sept 6) Silicates-Traces new MAN - part 2 (Oct28-29 G-probe)
ELEM: | Ba la | V ka | Mn ka | Fe ka | Co ka | Ni ka | Zn ka | Ca ka | P ka | Zr la | S ka | Cl ka | K ka | Ti ka | Cr ka | Na ka | Al ka | Mg ka | Si ka | Rb la | Sr la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 4 | 4 | 4 | 5 | 5 | 5 | 6 | 6 | 6
CRYST: | LiF | LiF | LiF | LiF | LiF | LiF | LiF | PET | PET | PET | PET | PET | PET | PET | PET | RAP | RAP | RAP | PET | PET | PET
ORDER: | 1 | 2 | 3 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3
ONTIM: | 135 | 115 | 139 | 25 | 115 | 115 | 115 | 20 | 67 | 100 | 92 | 65 | 40 | 125 | 224 | 25 | 65 | 293 | 65 | 170 | 154
Fall 2019 | CU Fe-Ti oxides 1.0 | [10 empty] | Full method description file finished
2019-09 (Sept 6) Silicates-Traces new MAN - part 4 OXIDES (Nov 26, Dec 12,13,14)
2019-09 (Sept 6) Silicates-Traces new MAN | [11 empty] | 2019-09-06 Silicates-Traces update MAN calibration.MDB | Detection limits from this file
ELEM: | V ka | Mn ka | Fe ka | Ni ka | Ca ka | Ti ka | Cr ka | Al ka | Mg ka | Si ka
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 2 | 2 | 3 | 4 | 4 | 5 | 5 | 6
CRYST: | LiF | LiF | LiF | LiF | PET | PET | PET | RAP | RAP | PET
ORDER: | 1 | 2 | 1 | 2 | 1 | 1 | 2 | 1 | 2 | 1
ONTIM: | 155 | 80 | 30 | 206 | 240 | 30 | 206 | 60 | 176 | 244
STDS: | 210 | 146 | 114 | 150 | 114 | 114 | 144 | 114 | 114 | 114
EDS 120 seconds (option for major-minor elements)
2019-09 (Sept 6) Silicates-Traces new MAN ELEM:: January 2021; 2019-09 (Sept 6) Silicates-Traces new MAN V ka: CU high-precision mafic glass with trace elements version 2.2 |  | Full method description file finished
Detection limits from this file
2019-09 (Sept 6) Silicates-Traces new MAN ELEM:: 2021-02 (Feb 8-10) NEW BASALT TRACES with Zn on RAP |  | 2021-02 Basalt Traces - NEW with Zn CLEAN NEW FILE.MDB
ELEM: | Ba la | V ka | Fe ka | Mn ka | Ni ka | Ca ka | P ka | Zr la | S ka | Cl ka | K ka | Ti ka | Cr ka | Na ka | Zn la | Mg ka | Si ka | Rb la | Sr la
BGD: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
BGDS: | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN | MAN
SPEC: | 1 | 1 | 2 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 4 | 4 | 4 | 5 | 5 | 5 | 6 | 6 | 6
CRYST: | LiF | LiF | LiF | LiF | LiF | PET | PET | PET | PET | PET | PET | PET | PET | RAP | RAP | RAP | PET | PET | PET
ORDER: | 1 | 2 | 1 | 2 | 3 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3
ONTIM: | 230 | 112 | 60 | 141 | 125 | 30 | 65 | 100 | 60 | 46 | 35 | 144 | 150 | 25 | 190 | 120 | 10 | 170 | 152
| 15 kV | 40 nA | EDS 10 seconds